# Investigating the performance of expert system in a multistage system for automatic detection of epileptic spikes

Anh-Dao Thi Nguyena,b, Duc-Tan Tranb, Nguyen Linh-Trungb

aUniversity of Technology and Logistics, Ho town, Thuan Thanh District, Bac Ninh, Vietnam

e-mail: daonta81@gmail.com

bVNU University of Engineering and Technology, 144 Xuan Thuy, Hanoi, Vietnam

e-mails: {tantd, linhtrung}@vnu.edu.vn

# Abstract

An automatic detection system for epileptic spikes is used as an assistive tool for epileptic analysis and diagnosis using electroencephalogram (EEG). Our system consists of four stages: preprocessing, feature extraction classification and expert system. In this paper, we concern to the performance of 4th stage (expert system) to choose the best threshold for detecting epileptic spikes efficiently. The proposed method is verified by experimental data acquired from ten patients (for training) and four patients (for testing).

Keywords: Epilepsy, EEG, spike, time-frequency, wavelet, artificial neural network, expert system.

1. Introduction

Epilepsy is a set of chronic neurological disorders. Epilepsy can be characterized by epileptiforms and seizures. Epileptiforms are waveforms that are related to epilepsy and it usually occurs before or after a seizure [1]. Epileptic seizures are obtained from abnormal, excessive or hyper synchronous neuronal activity in the brain. Automatic spike detection is very important because of providing important information such as spike density, patient syndrome. Over the last 40 years, good reviews on spike detection can be consulted in [2, 3, 4] and references therein.

The automatic detection systems for epileptic spikes can be categorized into two groups: simplex group, which is a simple structure, normally consisting of one or two engineering techniques [5,6] and multi-stage group which is built up by combining the multiple signal processing and machine learning techniques such as [7, 8, 9]. Usually, the multi-stage systems outperform the simplex systems in identifying epileptic spikes. Thus, our proposed automatic epileptic spike detection system consists of four stages: preprocessing, feature extraction, classification and expert system. Note that, the EEG signal should be put into the noise reduction process [10] before this multistage spike detection system. In this paper, we concern to the performance of 4th stage (expert system) to choose the most suitable threshold for the decision of the detecting epileptic spikes.

The paper is organized as follows. Section 2 describes our proposed automatic epileptic spike detection system. Experiments are presented in Section 3. Section 4 gives the results and discusses the results and Section 5 concludes the paper with some remarks for future works.

2. Theory

2.1. Proposed System Model

Figure 1 shows the block diagram of our proposed spike detection system which consists of 4 stages: preprocessing, feature extraction, classifier and expert system. First, all peaks of an EEG signal are automatically detected in 1st state (Preprocessing) and then, negligible peaks are identified and removed. Next, six features (the amplitudes, durations and slopes) [7], which associated to the morphology of an epileptic spike for each significant peak, are calculated and fed to three different perceptrons. Then, the peaks are separated into two groups of (1) non-epileptic spikes and 2) possible epileptic spikes.

Next, the possibility epileptic spikes are analyzed by continuous wavelet transform in the 2nd stage (Feature extraction). We then calculate seven wavelet features of a wavelet scale, which resembles an epileptic spike, from the 4th to the 8th scales [9]. Then, in the 3rd stage (Classifier), the extracted features are fed into the trained ANN. The spike score values, which are output of trained ANN, are in the range {0, 1}. If a value is closer to 1, the corresponding spike is likely to be an epileptic spike. Lastly, in the 4th stage (Expert system), we enforce a simple formula in parliamentary procedure to get rid of the pseudo-spikes which are settled about an epileptic spike.

Figure 1. Block diagram of the proposed multi-stage automatic spike detection system.

2.2. Expert system

In our system, an expert system is applied to reduce pseudo (false) spikes. By exploiting the spatial and temporal contextual information, the expert system analyzes EEG data which highly possible to be an epileptic spike. Specifically, in all channels, there exist delays of peak appearances. Epileptic events are usually presented by a “spike and slow-wave complex”. Because the slow-wave duration is around 150 - 350 ms therefore there exist no more than two spikes in an such time interval [11]. The expert system is contrived to do the following steps:

Step 1: Initiate a moving window with the length of 350 ms.

Step 2: Compute the weighting parameter for every spike existed in the window as ρ=0.5(a1+a2).

Step 3: Remove the spikes whose weight ρ fall below k ρmax, where 0<k<1, and ρmax is the local maximum of ρ within the windows of interest.

Step 4: Continue with the next window.

3. Experiments

3.1. Data Acquisition

In our research, EEG data is recorded by the EEG measurement system E-series of Compumedics. The system can record EEG data and movements of the patient. The proposed scheme was evaluated on EEG data recorded using an international standard 10 - 20 recording system with 19 channels. The sampling rate of record system is 256 Hz. The duration for each recording varies from 6 to 31 minutes.

Table 1. Training dataset

| Patient | Gender | Age | Duration | Neurologist |
| --- | --- | --- | --- | --- |
| 1 | Male | 7 | 10m 30s | 72 |
| 2 | Male | 13 | 9m 27s | 70 |
| 3 | Male | 15 | 9m 55 | 169 |
| 4 | Female | 10 | 11m 19s | 40 |
| 5 | Male | 18 | 10m 27s | 21 |
| 6 | Female | 7 | 22m 22s | 13 |
| 7 | Female | 18 | 30m 43s | 105 |
| 8 | Female | 17 | 22m 58s | 197 |
| 9 | Male | 11 | 9m 46s | 102 |
| 10 | Male | 13 | 17m 49s | 39 |

Table 2. Testing dataset

| Patient | Gender | Age | Duration | Neurologist |
| --- | --- | --- | --- | --- |
| 1 | Male | 8 | 13m 53s | 36 |
| 2 | Male | 11 | 9m 23s | 180 |
| 3 | Female | 10 | 9m 46s | 111 |
| 4 | Male | 6 | 5m 39s | 346 |

We carried out measurements on patients who were clinically diagnosed to suffer epilepsy. There are 14 epileptic patients in total, with 10 males and 4 females. The data from 10 patients were used for system training (Table 1) and the data from the other 4 patients were used for system evaluation (Table 2). The age of the patients ranges from 6 to 18. The spikes of the patient, which are labeled by the neurologist, are represented by parameter Neurologist in the tables in this paper.

3.2. Evaluation Metric

The sensitivity (SEN) is used as the metric for evaluating the operation of the scheme, which is delineated in [12] as follows

|  | (1) |
| --- | --- |

where the true positive (TP) is the number of spikes that both the system and the neurologist agree to be epileptic, and the false negative (FN) is the number of spikes that the neurologist labels as epileptic spikes but the system does not recognize so.

4. Results and Discussion

Tables 3-4 show the spike detection results from patients 1-4. The Table 5 summarizes the performance of the system for the best selection of the threshold (i.e. 0.4). The validation statistics of our systems for EEG data recorded in 4 epileptic patients are described in Tables 5. As we can see in Table 5, the sensitivity is very different among the patients (the smallest is 79.28% and largest is 91.33%). The difference in the value of sensitivity among the patients can be explained via the recording time (Duration) and the number of spike (Neurologist) in Table 2. For examples, patient 1 has 36 spikes labeled by the neurologist for a recording time of 13m 53s while the patient 5 has 346 spikes for 5m 39s.

Table 3. Spike detection results from patients 1-2

| Thresholds | Patient 1 | Patient 2 |
| --- | --- | --- |
|  | Neurologist | TP | FN | SEN (%) | Neurologist | TP | FN | SEN (%) |
| 0 | 36 | 29 | 7 | 80.55 | 180 | 146 | 34 | 81.11 |
| 0.1 | 36 | 29 | 7 | 80.55 | 180 | 146 | 34 | 81.11 |
| 0.2 | 36 | 29 | 7 | 80.55 | 180 | 146 | 34 | 81.11 |
| 0.3 | 36 | 29 | 7 | 80.55 | 180 | 146 | 34 | 81.11 |
| 0.4 | 36 | 29 | 7 | 80.55 | 180 | 146 | 34 | 81.11 |
| 0.5 | 36 | 29 | 7 | 80.55 | 180 | 146 | 34 | 81.11 |
| 0.6 | 36 | 29 | 7 | 80.55 | 180 | 146 | 34 | 81.11 |
| 0.7 | 36 | 28 | 8 | 77.78 | 180 | 146 | 34 | 81.11 |
| 0.8 | 36 | 25 | 11 | 69.44 | 180 | 146 | 34 | 81.11 |
| 0.9 | 36 | 24 | 12 | 66.66 | 180 | 138 | 42 | 76.66 |
| 1.0 | 36 | 23 | 13 | 63.88 | 180 | 119 | 61 | 66.11 |

Table 4. Spike detection results from patients 3-4

| Thresholds | Patient 3 | Patient 4 |
| --- | --- | --- |
|  | Neurologist | TP | FN | SEN (%) | Neurologist | TP | FN | SEN (%) |
| 0 | 111 | 88 | 23 | 79.28 | 346 | 318 | 28 | 91.90 |
| 0.1 | 111 | 88 | 23 | 79.28 | 346 | 318 | 28 | 91.90 |
| 0.2 | 111 | 88 | 23 | 79.28 | 346 | 318 | 28 | 91.90 |
| 0.3 | 111 | 88 | 23 | 79.28 | 346 | 318 | 28 | 91.90 |
| 0.4 | 111 | 88 | 23 | 79.28 | 346 | 317 | 29 | 91.61 |
| 0.5 | 111 | 88 | 23 | 79.28 | 346 | 314 | 32 | 90.75 |
| 0.6 | 111 | 88 | 23 | 79.28 | 346 | 306 | 40 | 88.43 |
| 0.7 | 111 | 87 | 24 | 78.38 | 346 | 280 | 66 | 80.92 |
| 0.8 | 111 | 82 | 29 | 73.87 | 346 | 246 | 100 | 71.09 |
| 0.9 | 111 | 74 | 37 | 66.67 | 346 | 207 | 139 | 59.82 |
| 1.0 | 111 | 64 | 47 | 57.66 | 346 | 169 | 177 | 48.84 |

Table 5. Performance results of the proposed system

| Patient | Gender | Age | Duration | Neurologist | TP | FN | SEN (%) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | Male | 8 | 13m 53s | 36 | 28 | 8 | 80.55 |
| 2 | Male | 11 | 9m 23s | 180 | 146 | 34 | 81.11 |
| 3 | Female | 10 | 9m 46s | 111 | 88 | 23 | 79.28 |
| 4 | Male | 6 | 5m 39s | 346 | 317 | 29 | 91.61% |

Figure 2: The relationship between the average of SEN and Threshold

Figure 2 shows the relationship between the average of SEN and Threshold. From Figure 2, we see that, threshold from 0 to 0.4, the value of SEN is constant and from 0.6 to 1.0, the SEN decreases to 0. Hence, in our system, the value of k is set to be 0.4 for the best performance of the expert system of all patients.

5. Conclusion

This paper proposed an investigation of the impact of the threshold on the performance of the fourth stage of the detection system for epileptic spikes. The range of the threshold from 0 to 0.4 will ensure the average sensitivity of 83.1%. This study has been successfully implemented and validated on real EEG data of epileptic patients. We also know that the epileptic spike amplitudes are not always stronger than the background EEG signal. In fact, sometimes epileptic spikes are not so distinguished due to its low voltage amplitude. Our expert system might make misjudgment in these cases. This is a limitation to the entire system and need to be investigated carefully in our future work.

6. Acknowledgments

The authors would like to thank Dr. Hoang Cam Tu from the Vietnam's National Hospital of Pediatrics for recording and interpreting the EEG data, and Mr. Dinh Van Viet from the Laboratory for Signals and Systems at the University of Engineering and Technology, Vietnam National University Hanoi, for assisting us to perform various simulations.

7. References

[1] K. Indiradevi, E. Elias, P. Sathidevi, S. Dinesh Nayak, K. Radhakrishnan, “A multi-level wavelet approach for automatic detection of epileptic spikes in the electroencephalogram,” Computers in Biology and Medicine, Vol. 8, No. 7, pp. 805-816, 2008.

[2] J. Gotman, “Automatic detection of seizures and spikes”, Journal of Clinical Neurophysiology, Vol.16, No. 2, pp.130-140, 1999.

[3] S. B. Wilson, R. Emerson, “Spike detection: A review and comparison of Algorithms”, Clinical Neurophysiology, Vol 113, No. 12, pp 1873-1881, 2002.

[4] A. T. Tzallas, M. G. Tsipouras, D. G. Tsalikakis, E. C. Karvounis, L. Astrakas, S. Konitsiotis, M. Tzaphlidou, “Automated epileptic seizure detection methods: A review study, Epilepsy-histological, electroencephalographic and psychological aspects”, edited by D. Stevanovic, pp. 75-98, 2012.

[5] G. Xu, J. Wang, Q. Zhang, S. Zhang, J. Zhu, “A spike detection method in EEG based on improved morphological Filter”, Computers in biology and medicine, Vol. 37, No. 11, 1647-1652, 2007.

[6] T.-W. Shen, X. Kuo, Y.-L. Hsin, “Ant K-Means Clustering Method on Epileptic Spike Detection”, in: Natural Computation, 2009. ICNC'09. Fifth International Conference on, Vol. 6, pp. 334-338, IEEE, 2009.

[7] N. Acir, Oztura, B. Baklan et al., “Automatic Detection of Epileptiform Events in EEG by a Three-Stage Procedure Based on Artificial Neural Networks”, IEEE Transactions on Biomedical Engineering, Vol. 52, No. 1, 30-40, 2005.

[8] H. Hamid, B. Boashash, “A time-frequency approach for EEG spike detection”, Iranica Journal of Energy & Environment, Vol. 2, No. 4, 390-395, 2011.

[9] H. S. Liu, T. Zhang, F. S. Yang, “A multistage, multimethod approach for automatic detection and classification of epileptiform EEG”, IEEE Transactions on Biomedical Engineering, Vol. 49, No. 12, 1557-1566, 2002.

[10] Anh-Dao, N. T., Duc-Nghia, T., Thi-Hao, N., Duc-Tan, T., Linh-Trung, N. , "An effective procedure for reducing EOG and EMG artefacts from EEG signals." Advanced Technologies for Communications (ATC), 2013 International Conference on. IEEE, 328-332, 2013.

[11] C. F. Boos, F. M. d. A. G. R. Scolaro, V. P. Maria do Carmo, “Automatic Detection of Paroxysms in EEG Signals using Morphological Descriptors and Artificial Neural Networks”, INTECH Open Access Publisher, 2011.

[12] Casson, A. J., E. Luna, E. Rodriguez-Villegas, “Performance metrics for the accurate characterization of interictal spike detection algorithms”, Journal of neuroscience methods, Vol. 177, No. 2, 479-487, 2009.